BED & SIDE RAIL SAFETY AUDIT

DDS Public Programs	 Southbury Training School

DDS Qualified Provider      _______________________________________________

(agency name)

Individual:      					Date of Review:

Address:

Bed Serial Number:      				Location of Serial Number:

Type of Bed (check one):

Fully electric hospital bed 		 Semi electric hospital bed		 Manual hospital bed

Standard bed with added equipment 	 Other:

Circle the diagram that best indicates the Rail Configuration on the bed being assessed.

- Bed Operations: Remove all sheets, blankets, bolsters, air mattresses, etc., before proceeding

- Semi-electric or manual beds:

Cranks smoothly		Yes  No  Comment:

Head/foot/gatch		Yes  No  Comment:

Trendelenburg			Yes  No  Comment:

Hi-Lo				Yes  No  Comment:

Manual crank backup		Yes  No  Comment:

- Fully electric beds:

Cranks smoothly		Yes  No  Comment:

Head/foot/gatch		Yes  No  Comment:

Trendelenburg			Yes  No  Comment:

Hi-Lo				Yes  No  Comment:

- Semi-electric or fully electric beds:

Cords/plugs free of damage		Yes  No  Comment:

Cords secured properly		Yes  No  Comment:

Controls work properly		Yes  No  Comment:

(e.g., buttons/switches)

Grounded prong plug on bed		Yes  No  Comment:

Grounded wall outlet available	Yes  No  Comment:

- All beds

Wheels operate smoothly	Yes  No  Comment:

Wheel locks work		Yes  No  Comment:

Headboard secured properly	Yes  No  Comment:

Footboard secured properly	Yes  No  Comment:

Rust				Yes  No  Comment:

Dents				Yes  No  Comment:

Sharp edges			Yes  No  Comment:

Finish is intact			Yes  No  Comment:

Surface is smooth		Yes  No  Comment:

- Mattress:

- General:

Type

Intact				Yes  No  Comment:

Stains				Yes  No  Comment:

Odor				Yes  No  Comment:

Fits properly			Yes  No  Comment:

Is there a gel overlay 		Yes  No  Comment:

Other overlay	 		Yes  No  Comment:

Replace all sheets, blankets, bolsters, air mattresses, etc., before proceeding.

- Observe and identify equipment present:

To keep mattress up and in place (top to bottom):

Bracket			Yes  No  Comment:

In proper position?	Yes  No  Comment:

Blankets			Yes  No  Comment:

Bolsters			Yes  No  Comment:

Pillows			Yes  No  Comment:

Other			Yes  No  Comment:

To fill in gaps between mattress and side rails (Zone 3 – gap to measure 2 inches or less)

Blankets			Yes  No  Comment:

Bolsters			Yes  No  Comment:

Pillows			Yes  No  Comment:

Other			Yes  No  Comment:

Measurements: (right or left side of the bed is identified by standing at the head of the bed and facing the footboard)

Gap (if any) between mattress and each side rail (Zone 3):

Left rail       / Right rail       (If greater than 2 inches then bolster is required)

Gap (if any) between mattress and headboard and/or footboard (Zone 7):

Headboard       / Footboard       (Assess need for bolster if greater than 3 inches)

Distance from top of mattress to top of bed rail:

(Please note, there is no standard measurement for the distance from the top of the mattress to the top of the bed rail.  However, it is important to understand that whenever the mattress that came with the bed is replaced with a different mattress or other equipment such as an air mattress, foam, egg crate, etc. designed to be placed on top of the mattress, this raises the level of the individual lying on the bed and may pose a risk of the individual falling over the bed rail.)

Is the height of the mattress alone or with any additional equipment lower than the top of the rail and low enough to prevent the individual from falling over the rail?    Yes  No

If no, specify what measures have been taken to correct this issue:

- Bed Side Rails:

Present					Yes  No  Comment:

Type				        Single  Full  Split rail

Side rails and bed from same		Yes  No  Unknown  Comment:

manufacturer

If no, are they compatible		Yes  No  Comment:

Side rails operate correctly		Yes  No  Comment:

Play detected (loose)		Yes  No  Comment:

Bowed or warped			Yes  No  Comment:

Side rail pads (if present):

Intact				Yes  No  Comment:

Stains				Yes  No  Comment:

Odor				Yes  No  Comment:

Fits properly			Yes  No  Comment:

Measurements:

Equal distance headboard to 	Yes  No  Comment:       (Zone 6)

footboard				(If greater than 3 inches then action needed)

Equal Distance side to side	Yes  No  Comment:       (Zone 3)

Distance within the rails (Zone 1)

Left rail       / Right rail       (If greater than 4¾ inches then action needed)

Distance between the mattress and the base of the rails with HOB elevated to approximately 30 degrees (Zone 2)

Left rail       / Right rail       (If greater than 3 inches then action needed)

Distance between the split between the hemi rails (Zone 5)

Left side       Right side       (If greater than 12.5 inches then action needed)

Conclusion/Recommendations:

The safety and operating conditions of this bed, side rails, and related equipment is consistent with FDA guidance to reduce entrapment in bedrails. Yes  No  Comment:

https://www.fda.gov/medical-devices/consumer-products/bed-rail-safety

https://www.fda.gov/medical-devices/bed-rail-safety/recommendations-health-care-providers-about-bed-rails

The following changes are required to make the bed and side rails consistent with FDA guidance standards to reduce entrapment in bedrails.

Audit team members:

| First Name | Last Name | Date |
| --- | --- | --- |

(The following review is to be completed only if entrapment risks are identified)

Follow-up review of 	Bed and Side Rail Safety Audit.

Date of Review:

The audit team members have reviewed the changes made to this bed, as outlined in the recommendations above. All recommendations have been satisfactorily completed.

Audit team members:

| First Name | Last Name | Date |
| --- | --- | --- |

See Appendix A: Potential Entrapment Zones as Identified by FDA Guidance below.

Zone 3 Entrapment between the rail and the mattress

Zone 6 Entrapment between the end of the rail and the side edge of the head or foot board

Zone 7 Entrapment between head or foot board and the mattress end

Zone 6 Entrapment between the end of the rail and the side edge of the head or foot board

Zone 3 Entrapment between the rail and the mattress

Zone 2 Entrapment under the rail, between the rail supports or next to a single rail support

Zone 6 Entrapment between the end of the rail and the side edge of the head or foot board

Zone 7 Entrapment between head or foot board and the mattress end

Zone 6 Entrapment between the end of the rail and the side edge of the head or foot board

Appendix A

Potential Entrapment Zones

as Identified by FDA Guidance Standards

Potential Entrapment Zones as Identified by FDA Guidance Standards

Zone 1 – Entrapment within the rail	Zone 2 – Entrapment under the rail, between the rail supports or next to a single

rail support

Zone 3 – Entrapment between the rail and the mattress

Zone 4 – Entrapment under the rail, at end of rail

Zone 5 – Entrapment between split bed rails

Zone 6 – Entrapment between the end of the rail and the side edge of the head or foot board

Zone 7 – Entrapment between head or foot board and the mattress end