**BED & SIDE RAIL SAFETY CHECKLIST**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **Date:**      **\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**  **Name:**      **\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**  **Address:**      **\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_** | **Condition:**  **S – Satisfactory**  **U – Unsatisfactory, needs repair** | **Indicate action taken, if unsatisfactory** |  | **Condition:**  **S – Satisfactory**  **U – Unsatisfactory, needs repair** | **Indicate action taken, if unsatisfactory** |
| **Bedframes and Rails** |  |  | **Mattress** |  |  |
| Clean | **S**   **U** |  | Fits properly | **S**   **U** |  |
| Rust free | **S**   **U** |  | Clean | **S**   **U** |  |
| No dents | **S**   **U** |  | No cracks/holes/rips | **S**   **U** |  |
| No sharp edges | **S**   **U** |  | No stains/odor | **S**   **U** |  |
| Mattress guard present | **S**   **U** |  | Gap between mattress and each side rail no more than 2 inches | **S**   **U** |  |
| Mattress guard in proper position | **S**   **U** |  |
| **Wooden Beds** |  |  | Headboard to mattress gap not more than 3 inches with bed flat | **S**   **U** |  |
| Finish Intact and Smooth | **S**   **U** |  |
| **Semi-Electric and Fully Electric Beds** |  |  | Footboard to mattress no more than 3 inches | **S**   **U** |  |
| Cord/plug undamaged | **S**   **U** |  |
| Cord secured correctly | **S**   **U** |  | Gap between mattress and base of rails with HOB elevated 30 degrees no more than 3 inches | **S**   **U** |  |
| Controls work correctly | **S**   **U** |  |
| Grounded outlet used | **S**   **U** |  |
| **Bed Operations** |  |  | Distance between split in split rail configurations not more than 12.5 inches | **S**   **U** |  |
| Cranks smoothly | **S**   **U** |  | **Egg crate foam/**  **Air Mattress** |  |  |
| Wheels operate smoothly | **S**   **U** |  | Clean | **S**   **U** |  |
| Wheel lock works | **S**   **U** |  | No cracks/holes/rips | **S**   **U** |  |
| Headboard secured | **S**   **U** |  | No stains/odors | **S**   **U** |  |
| Footboard secured | **S**   **U** |  |  |  |  |
|  |  |  |  |  |  |

**BED & SIDE RAIL SAFETY CHECKLIST**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **Date:**      **\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**  **Name:**      **\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**  **Address:**      **\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_** | **Condition:**  **S – Satisfactory**  **U – Unsatisfactory, needs repair** | **Indicate action taken, if unsatisfactory** |  | **Condition:**  **S – Satisfactory**  **U – Unsatisfactory, needs repair** | **Indicate action taken, if unsatisfactory** |
| **Side Rails** |  |  | **Side Rail Pads** |  |  |
| No play detected | **S**   **U** |  | Clean | **S**   **U** |  |
| Operate correctly | **S**   **U** |  | No cracks/holes/rips | **S**   **U** |  |
| Not bowed or warped | **S**   **U** |  | No stains/odor | **S**   **U** |  |
| Distance b/w mattress & headboard same as distance b/w mattress & footboard? | **S**   **U** |  |  |  |  |
| Distance b/w mattress & side rail on Right side same as distance b/w mattress & side rail on Left side? | **S**   **U** |  |  |  |  |
|  | | | | | |
| **Name and Signature of Person Completing Checklist:**      **\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_** | | | | | |